NIH Image as a limited nuclear medicine view stationMark D. Wittry, James W. Fletcher, J. Stephen Farris, and James L. Daly.Saint Louis University Health Sciences Center, Saint Louis, MissouriImage is a Macintosh-based software platform that has been developed by the NationalInstitutes of Health to support image manipulation. Image is a public domainapplication that is provided to any prospective user at no charge and is available on avariety of bulletin board services or directly by InterNet. A variety of imagemanipulation tools have been hard-coded into Image that have direct utility fornuclear medicine applications. However, these tools have not been optimized foreasy use. The Division of Nuclear Medicine at Saint Louis University HealthSciences Center has developed a structured series of image manipulation routines(known as �Image Macros�) designed specifically for nuclear medicine with thePascal-based macro compiler supplied with Image.Three levels of image manipulation are supported within �Image Macros�: 1. datatranslation from the vendor�s proprietary format to one readily recognized withinImage, 2. basic image manipulation tools that work on files of either single imagesor a multi-image set (called a �stack�), and 3. simple processing macros.Data Translation:The Department of Nuclear Medicine at Saint Louis University Health Sciences Centerconsists of hardware from 6 vendors with software in 8 data formats on 4 differentoperating systems. A thin-wire EtherNet backbone with TCP/IP and AppleShare transferprotocols allow bi- directional data transfer from all but one vendor (data may bereceived from but not transmitted to Elscint X09 terminals). Image data istransferred to Macintosh platforms (ranging from the IIsi up to the Quadra 950)for translation into Image. Translator modules have been written into �Image Macros�that translate the vendor�s proprietary data formats into a common format thatImage can easily load for display and analysis (data formats currently supportedinto Image include Siemens microDELTA, Siemens ICON, CTI ECAT,ADAC Pegasys and Elscint APEX). The common format is a modified version ofthe Siemens microDELTA format, chosen only because of the principle author�s extensivefamiliarity with this format. The matrix, bits per pixel, maximal pixel value, calibrationfactor for pixel units (e.g. counts/pixel/second), number of images and patient�sidentifying data are automatically incorporated into Image for use in both image displayand quantitation. The slow speed of the non-compiled �Image Macros� on a 680x0 CPUcan severely limit data-access speed for large data sets (e.g. a PET study consisting of310 128 x 128 x 16 bit images) and has required the initial translation of ACR-NEMA v2.0and ECAT v6.5 data to be performed outside of Image by stand-alone applicationsdeveloped with Symantec�s Think C, version 6.0. Translated image sets quickly load intoImage except for ECAT data from CTI where a 128 x 128 x 16 bit image takes 6 secondsa slice on a 68030-based Macintosh CPU.Image Manipulation:The user programming environment contained within Image consists of calls tobasic image manipulation tools such as image rotation and animation. Examplemacros are provided within the �Macros� folder. Common image manipulationtools have been optimized into a pre-loaded series of macros that works bothon individual images as well as �stacks� of images (i.e dynamic and multi-imageSPECT studies). These routines are defined in a file called �Image Macros� thatis automatically loaded at the launch of Image. All common image manipulationmacros are quickly accessible through single keystrokes based upon the logicof the command (�V� - flip the image in the Vertical direction, �A� - Animate allwindows). The �Image Macros� are also accessible from the �Special� pull-downmenu in the Image menubar. Currently supported macros with a brief explanationof the function of each one are listed below:macro '[A]    Animate All'     Provides a seamless cine loop of all open image windows					provided that each window is a �stack� of an equal number					of images.macro '[B]    Back projection'					Performs SPECT back projection. The user must enter the					start angle in degrees and length of SPECT arc in degrees.					Clockwise rotation is [+], counter-clockwise [-]. The design of					of this macro performs a 'pseudo' ramp filtration through     image rescaling. The projection data should undergo a					2-dimensional filtration prior to back projection.macro '[C]   Crop images'     Extracts the area of an image contained within a region of					interest into a separate window (this will extract the					corresponding area from all slices in a �stack� of images).macro '[E]    Translate Elscint Images'					Converts Elscint APEX� images into a TIFF format for NIH Image.macro '[F]    Filter'    	Perform one of 4 filter operations on the image(s) in the	foreground window. 							1: a 9 point spatial smooth							2: a 'sharpening' filter							3: a noise reduction filter							4: a sharp metz filtermacro '[G]    Gray Scale'					Resets the LookUp Table to 'gray scale'.macro '[H]    Flip Horizontal'     Flips the image(s) in the horizontal plane (around the �y� axis).macro '[I]    Invert LUT'       Inverts the LookUp table for the foreground image.macro '[M]    Multiply by Constant'       Multiplies all pixels in a data set by a user defined constant (useful for       ennhancing low count projection data prior to filtration and back projection).macro '[O]    Oblique Angulation'      Variable rotation of the images (useful for the	reconstruction of �Oblique�      images from the SPECT	transaxial data set).macro '[P]    Translate Pegasys Images'						Converts ADAC Pegasys� images into a TIFF format for NIH Image.macro '[R]     Rotate Images'						L: Rotates the image(s) in the foreground window 90� to the	left (counterclockwise).						R: Rotates the image(s) in the foreground window 90� to the	right (clockwise).macro '[S]    Shuffle Images'       Resorts multihead acquisitions (SPECT or dual-energy settings)							so that all images from head #1 are in sequence and followed by       all images form head #2 etc....macro '[T]    Translate MicroDELTA Images'       Converts Siemens MicroDELTA� images into a TIFF format for NIH Image.macro '[V]    Flip Vertical'       Flips the image(s) in the vertical plane (around the �x�	axis).macro '[Z]     Zoom images�       Zooms the matrix of the image(s) in the foreground 	window by avariable factor							while performing a	bi-directional pixel interpolation.macro 'Invert Images'      Inverts the density of the pixel values for the image set (0->255, 255->0).macro 'Dispose all...'       Closes all open image windows. Does not close text windows.Simple Processing Macros:Image contains a �toolbar� of commonly used image interrogation tools such asrectangular, circular and freehand regions of interest. Multi-step macros havebeen created to calculate the left ventricular ejection fraction using a 2 framemanual method as well as macros to slice a SPECT reconstruction data set intoeach of the conventional planes (transaxial, sagittal, coronal, and oblique).Left Ventricular Ejection Fraction:macro 'Window for EF Calculation'					The window containing the �best septal� or left anterior					oblique view from a radionuclide ventriculogram should be					in the foreground. This macro, when launched, will create					a text window prompting all subsequent steps as well as					displaying the computed values from each step. The first					step consists of a prompt to draw a rectangular region of					interest (any type of a region of interest will suffice) over					the left ventricle in the slice (image) displayed. Once this					region is created, the user is instructed to proceed to the					second macro entitled �Identify ED (end-diastole) and ES					(end-systole).macro 'Identify ED and ES'     This macro interrogates the region of interest drawn by					the user to calculate the total counts contained within the					region for each �slice� in the image �stack�. The values are					displayed in the text window. The macro takes the slice					with maximum counts and creates a new single image					window called �ED�. Similarly, the macro takes the slice					with minimal counts and duplicates single image windows					called �ES� and �Bkg�. The user is prompted to draw					�freehand� regions of interest on each window					(ED: end-diastole, ES: end-systole, Bkg: background)					then continue	to the �Calculate LVEF� macro.					Note: The algorithm of this macro is totally flexible. Slice					1 or 2 is not assumed to automatically be end-diastole.macro 'Calculate LVEF'    This macros interrogates the 3 regions drawn to compute					a simple 2 frame left ventricular ejection fraction with its					value, the corresponding slice numbers, the size and					counts of each region displayed. Time activity curves are					not supported at this time.macro 'Transaxial to Coronal...�     This macro reconstructs the coronal slices from a					transaxial data set. The entire matrix will be					reconstructed if no region of interest is used to define					limits. Only the area defined will be reconstructed if a					rectangular region of interest is drawn on the slice in the					foreground window prior to selecting this macro.macro 'Transaxial to Sagittal...'    This macro reconstructs the sagittal slices from a					transaxial data set. The entire matrix will be					reconstructed if no region of interest is used to define					limits. Only the area defined will be reconstructed if a					rectangular region of interest is drawn on the slice in the					foreground window

prior to selecting this macro.Short axis slices may be created from SPECT myocardial perfusiontransaxial data sets by performing the following sequence of macros:macro '[O]    Oblique Angulation'    Use the �angle� tool from the menubar to define the angle					that the apex is offset from the �y� axis (displayed in the					�Info� window in the lower left hand corner of the					Macintosh display screen). Enter the number of degrees to					rotate when prompted.macro 'Transaxial to Sagittal...'     Use the rectangular region of interest tool to define the					limits of the sagittal reconstruction. Be generous!macro '[O]    Oblique Angulation'    Use the �angle� tool from the menubar to define the angle					that the apex is offset from the �y� axis (displayed in the					�Info� window in the lower left hand corner of the					Macintosh display screen). Enter the number of degrees to					rotate when prompted. This realigns the sagittal images					into �true� vertical long axis slices.macro 'Transaxial to Coronal...�    Use the rectangular region of interest tool to define the					limits of the short axis reconstruction. macro '[H]    Flip Horizontal'     The displayed images will be inverted (septum to the right,					lateral wall to the left). This command reorients the slices					into the ACC/SNM convention.					Note: The �vertical long axis� slice may be reoriented to					the ACC/SNM convention by bringing this window to the					foreground and selecting the macro �Rotate Right�.A macro to allow calculation of the Standardized  Uptake Value for PETwhole body oncology studies as well as designing �time activity� displaycurves for dynamic studies are being developed.NIH Image is freeware available from most electronic bulletin board services.The macros created by Saint Louis University Health Sciences Center,Division of Nuclear Medicine are available as freeware directly from:Dr. Mark D. Wittry, M.D., F.A.C.C.Department of Nuclear MedicineSaint Louis University Health Sciences Center3635 Vista Avenue at Grand Blvd.P.O. Box 15250Saint Louis, Missouri, 63110-0250.Inquiries may be directed to Dr. Wittry at (314) 577-8047, (314) 268-5144 (FAX).Please submit a blank, Macintosh formatted 3.5� floppy disk and enclose $5.00for shipping and handling. We hope to upload the latest version of NIH Imageand the image macros to the Lunis bulletin board in the near future.Please note: Saint Louis University and Dr. Wittry make no warranty orrepresentation, either expressed or implied, with respect to this software,its quality, performance, merchantability, or fitness for a particular purpose.This software is provided �as is�, and you, the user, must assume and acceptthe entire risk as to its quality and performance. No responsibility will be taken by Dr. Wittry or Saint Louis University for anydiagnosis or interpretation of results produced from  this software packagewhatsoever.